# ‏‏

# SCIENTIFIC REPORT II

# Submitted to: Alpha Bio Tec.

# MODELLING INSERTION OF TWO DIFFERENT DENTAL IMPLANTS

Prepared by:

Prof. D. Rittel, Dr. A. Dorogoy and Dr. K. Shemtov-Yona

Faculty of Mechanical Engineering

Technion

September 2017

For the elimination of any doubt, it is hereby stressed that the staff member and/or the Technion and/or the Technion Research and Development Foundation Ltd. will not be liable for any property damage and/or corporal damage and/or expense and/or loss of any kind or sort that will be caused or may be caused to you or anyone acting on your behalf, in consequence of this statement of opinion or this report, or in any connection to it.

- Numerical insertion of Alpha-Bio dental implants

The process of inserting two commercial implants  (ICE_3.75_11.5_1011 , http://alpha-bio.net/il/products/implant-system/internal-hex-implants/neo/, figure 1)  into the mandible bone, was simulated  numerically using  the commercial finite elements package Abaqus Explicit version 6.14-2 (Simulia, 2014a). The model is three dimensional and the analyses are dynamic, continuous and non-linear. The simulation assumed a rigid implant rotating at a constant rotational velocity with an axially applied load downward. The initial insertion cavity is shown in Fig. 2 and is 11.5 mm in depth. The applied rotational velocity was 120 rpm, while the vertical load was (arbitrarily) set to 5 N. Note that the commonly applied angular velocities are close to 30 rpm, so that the above-quoted higher values used are used for the sake of computational time reduction only. Consequently, bone heating effects related to the insertion process are not considered in this work, even if they could be quite significant at high insertion speeds (Lee et al., 2012). The dimensions, material models and parameters, mesh properties and boundary conditions are detailed in the sequel.

It is important to note that this report is a preliminary report aimed among other things at assessing the suitability of the finite element code Abaqus for Alpha-Bio purposes. Future analyses will require Alpha Bio to purchase their own license.

  - Assembly dimensions

The sketches of the implants were provided by Alpha-Bio. The implants are 3.75 mm in diameter, 11.5 mm in length. They have conical shape and threads that increase in thickness along the implant body, as  shown in Fig. 1a,b. Implant 2 (Fig. 1b) has an upper groove which does not exist in Implant 1.

| a. | b. |
| --- | --- |

Figure 1: The Alpha-Bio implants: 3.75mm diameter, 11.5mm length.

- Implant 1. b. Implant 2.

Figure 2: The Alpha-Bio osteotomy  for the implants shown in Fig. 1.

An initial osteotomy having a maximum diameter of 3.65 mm and a total length of 11.5 mm (Fig. 2) is cut on the  upper bone face as shown in the cut views of Figs. 3d,e. The bone is modeled by  extruding a typical cross section of the mandible at the facial region 20 mm along the Z axis. It consists of cortical bone (green in Fig 3) having a thickness of ~3 mm, and of cancellous bone (grey in Fig. 3). The isometric view of the mandible bone and implant is shown in Fig. 3a.  Side views of Fig. 3a are shown in Figs. 3b and 3c. Cut views  and dimensions are shown in Figs. 3d and 3e.

| a. |  |
| --- | --- |
| b. | c. |
| d. | e. |

Figure 3: Assembly of the implant 2 and bone and boundary conditions at the beginning of the insertion process. a. Isometric view of the meshed assembly. b. X-Y side view. .c. Z-Y side view. d. X-Y cut view.  e. Z-Y cut view. All dimensions are in mm.

Material properties

The implant was modeled as a rigid (non deformable) body. The reason for this assumption is the reduction of computational time along with the fact that there was no request to analyze the stress and strain fields in the implant itself. The bone constituents (trabecular and cortical) were assumed to be perfectly bonded.  The material properties of the bone constituents are the same as used in Dorogoy et al. 2017, and are briefly detailed  in the sequel.

Cortical bone

The cortical bone was modeled as an elastic-(almost) perfectly plastic material with Drucker-Prager (DP) pressure-sensitive behavior.  Ductile failure with damage evolution (Simulia, 2014b) was used as a failure criterion. The tensile and compressive yield strengths were taken as 120  and 180 MPa, respectively (Fig. 4). The plastic strain at fracture was set to 1%. in tension and 2% in compression (Keaveny and Hayes, 1992). The “damage evolution value”, was set to zero assuming abrupt  failure of the cortical bone.

Figure 4: Unidirectional tension and compression of a cortical bone. (reprinted from Dorogoy et al 2017).

Cancellous bone

The cancellous bone is a cellular material, and is approximated here by an elastic-(almost ideal) plastic material model. The Young’s modulus, Poisson ratio and fracture strain were all taken from Guan et al. (Guan et al., 2011), while “representative” yield stresses were  set to vary in the range of 20 < σY < 47 MPa.  Using in the numerical model σY = 47 MPa corresponds approximately to bone volume of ~47%, while σY = 20 MPa corresponds to a bone volume fraction of 27%-31% (Hernandez et al., 2001; Hambli, 2013). Let us remind here that Bone Volume Fraction (BV/TV) is defined as the volume of mineralized bone per unit volume of the sample.

Numerical experimentation with different failure strains together with the elastic-plastic model revealed that the value of the failure strain has a small effect on the torque results, hence the value of εpf=0.135 (Guam et al, 2011) was used. The trabecular bone is softer than the cortical bone, hence a small value of 10 µm was given as damage evolution. This value was given just to increase stability of the numerical solution during element deletion and does not affect the torque results.

The typical material properties, based on (Guan et al., 2011; Reilly and Burstein, 1974; Keaveny and Hayes, 1992; van Staden et al., 2008), were used in the simulations and are summarized in Table 1.

|  | ρ [Kg/m3] density | E [MPa] Young’s modulus | ν Poisson ratio | σY [MPa] yield stress | β [o] Drucker Prager | εpT [%] Fracture plastic strain tension | εpC [%] Fracture plastic strain compression | Damage evolution [μm] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cortical bone | 1900 | 18000 | 0.35 | 180 | 30 | 1 | 2 | 0.0 |
| Cancellous bone | 1000 | 700 | 0.35 | 20-32-47 | 0 | 0.135 | 0.135 | 10 |

Table 1: Material properties  (Dorogoy et al., 2017).

2.3 Mesh

The meshed assembly is shown in Fig. 5a. Implant 1 (for example) was meshed with 15439 rigid shell elements. 14637 linear quadrilateral elements of type R3D4 and 802 linear triangular elements of type R3D3.  The meshed implant is shown in Fig. 5b.

The bone was meshed with 2684573  linear tetrahedral elements of type C3D4. The region around the cavity was meshed with small elements of size ~0.1 mm as shown in Fig5c.

| a. |  |
| --- | --- |
| RP |  |
| b. | c. |

Figure 5: a. The meshed assembly. b. The meshed implant !.   c. The exposed mesh near the cavity. RP (Reference Point) indicates the point of application of vertical load and implant rotations.

2.4  Boundary conditions

Symmetry conditions were applied on both planes of the bone with the normal pointing in the Z direction (nz = ±1).  A line along on the plane nz = ±1 (Fig. 3a) was fixed in all directions to prevent  rigid body motions of the assembly.

The rigid implant was only allowed to rotate and move downwards in the y direction. A constant angular velocity (ωy) of 120 rpm was applied to the implant on the reference point shown in Fig. 6a.  A constant downward force (F), of 5N was applied as well on the rigid implant on the reference point. The angular velocity and vertical load are shown schematically in Fig. 3a.

The general contact algorithm of Abaqus (Simulia, 2014b) was used with element-based surfaces which can adapt to the exposed surfaces of the current non-failed elements. Abaqus' frictional tangential behavior with the penalty formulation was adopted. A frictional Coulomb contact was used with a constant coefficient of friction 0.61 (Grant et al., 2007; Guan et al., 2011). All the surfaces that may become exposed during the analysis, including faces that are originally in the interior of bone were included in the contact model. We assumed that contact nodes still take part in the contact calculations even after all of the surrounding elements have failed. These nodes act as free-floating point masses that can experience contact with the active contact faces (Simulia, 2014b).

- Results

A cut view showing the assembly of the implant and the bone at time zero where the insertion depth is defined as zero is shown in Fig. 6a. The implant is inserted initially into the osteotomy a depth of ~4 mm before application of load. At time zero there is no contact between the implant and the inner walls of the osteotomy. The insertion depth from the initial state shown in Fig. 6a, versus time is shown in Fig. 6b. The results of Fig. 6b is for a trabecular bone having a yield stress of 32 MPa.  The   insertion depths of   d = 2, 4, 6 and 7.3 mm are marked in Fig. 6b.

The location of the implant within the bone, Mises stress distribution in the bone and plastic strain distribution in the bone at the locations marked in Fig. 6b are shown in Fig. 7.  Figures 7a-c are for d = 2mm. Figures 7d-f are for d = 4 mm. Figures 7g-i are for d = 6 mm. Figures 7j-l are for d = 7.3 mm, when the implant is fully inserted into its cavity.

a.

b.

Figure 6: a. A cut view showing the assembly of the implant and the bone at time zero where the insertion depth is defined as zero. b. The insertion depth versus time for a trabecular bone having a yield stress of 32 MPa. The  insertion depths of  d = 2, 4, 6 and 7.3 mm  are marked.

| d = 2.0 mm d = 2.0 mm |  |  |
| --- | --- | --- |
| d = 4.0 mm d = 4.0 mm a. | b. | c. |
| d. | e. | f. |
| d = 6.0 mm d = 6.0 mm |  |  |
| g. | h. | i. |
| d = 7.3 mm d = 7.3 mm |  |  |
| j. | k. | l. |

Figure 7: The location of the implant within the bone, Mises stress distribution in the bone and plastic strain distribution in the bone at the locations marked in Fig. 6b.

The resisting insertion torque (IT) evolutions for trabecular bone having σy =  20, 32  and 47 MPa are shown in Fig. 8. The IT due to the insertion of implant 1 is shown in Fig. 8a while the IT due to insertion of implant 2 is shown in Fig. 8b.

a.

b.

Figure 8: Insertion torque for trabecular bone having σy = 20, 32 and 47 MPa.

a. Insertion of implant 1. b. Insertion of implant 2.

A comparison of the IT evolution due to the insertion of implant 1 and implant 2 is shown in Fig. 9. The IT is plotted against the insertion depth. The comparison is for trabecular bone having σy =  20, 32 and 47 MPa. Only the filtered values are shown for clarity reasons.

Figure 9: Comparison of the insertion torque due to the insertion of implant 1 and implant 2 for trabecular bone having σy =  20, 32 and 47 MPa.

Effect of the cortical bone on the insertion torque

The insertion of both implants into a bone made of  cortical bone only was modeled. The meshed assembly is shown in Fig. 10a.  The exposed bone is shown in Fig. 10b.

| a. | b. |
| --- | --- |
| c. |  |

Figure 10: a. The meshed assembly. b. The exposed bone. c. The exposed insertion region.

The exposed insertion region is shown in Fig. 10c.

The resisting torque during the insertion is shown in Fig. 11. The position of the implant at full insertion (d = 7.68 mm) is shown by the insert to the right. It can be observed that by further inserting the implants the resisting torque become smaller and reach zero  for both implants at d = 10 mm.

Figure 12 is a combination of Fig. 9 and Fig. 11. It shows the contribution of the cortical bone alone (Fig. 12)  to the  torque in comparison to the torques of the full bones (Fig. 10) which are made of trabecular and cortical phases.

Figure 11: The insertion torque  of implant 1 and implant 2. The position of the implant at d = 0 mm and d = 7.68 mm  is shown by the inserts.

Figure 12: A comparison between the torque carried by the cortical bone alone to the total torque which is carried by both the cortical and trabecular bone.

- Preliminary conclusions

- Two different implant geometries were modelled numerically and the evolution of the bone mechanical parameters were characterized.

- Since the two considered implants are quite similar in their geometry, with the exception of a small groove in implant 2, it can be observed that overall, the calculated mechanical parameters are relatively similar as well. This is expectable.

- While the groove of implant 2 has little mechanical effect, its presence may be beneficial from a clinical point of view, not addressed in this study.

- The trabecular bone is machined first, followed by the cortical, as shown in Fig 7. This is the result of the combined implant- osteotomy geometry.

- As shown in Fig. 12, beyond an insertion depth of ca. 6 mm, the trabecular bone is essentially responsible for the resultant torque.

- The cortical bone causes an insertion torque mostly for d=3-6 mm, as shown in Fig. 12.

- The first contact between the implant and osteotomy is at d = ~1.2 mm (Fig. 6b). The location of this contact within the bone is close to the interface  between the trabecular and cortical bone. At d = 7.3 mm (Fig. 7j) the implants are fully inserted within the bone. The sequence of contact (threads are cutting through the bone)  vs. d is shown in Figs. 7b, 7e, 7h and 7k  (see also attached video).

- The contribution of the cortical bone to the torque is less than 35 Nmm for all stages of the insertion.

- This preliminary study shows that implant insertion, to start, can be successfully simulated, paving the way for implant geometry and osteotomy optimization.

References

Dorogoy, A., Rittel, D., Shemtov-Yona, K., & Korabi, R. (2017). Modeling dental implant insertion. Journal of the Mechanical Behavior of Biomedical Materials, 68, 42-50.

Grant, J.A., Bishop, N.E., Götzen, N., Sprecher, C., Honl, M., Morlock, M.M., 2007. Artificial composite bone as a model of human trabecular bone: the implant-bone interface. J. Biomech. 40, 1158–64. doi:10.1016/j.jbiomech.2006.04.007

Guan, H., van Staden, R.C., Johnson, N.W., Loo, Y.-C., 2011. Dynamic modelling and simulation of dental implant insertion process—A finite element study. Finite Elem. Anal. Des. 47, 886–897. doi:10.1016/j.finel.2011.03.005

Hambli, R., 2013. Micro-CT finite element model and experimental validation of trabecular bone damage and fracture. Bone 56, 363–374. doi:10.1016/j.bone.2013.06.028

Hernandez, C.., Beaupré, G.., Keller, T.., Carter, D.., 2001. The influence of bone volume fraction and ash fraction on bone strength and modulus. Bone 29, 74–78. doi:10.1016/S8756-3282(01)00467-7

Keaveny, T.M., Hayes, W.C., 1992. Mechanical properties of cortical and trabecular

bone. Bone, 285–344.

Lee, J., Ozdoganlar, O.B., Rabin, Y., 2012. An experimental investigation on thermal exposure during bone drilling. Med. Eng. Phys. 34, 1510–1520. doi:10.1016/j.medengphy.2012.03.002

Reilly, D.T., Burstein, A.H., 1974. The mechanical properties of cortical bone. J Bone Jt. Surg Am 56, 1001–1022.

Simulia, 2014a. Abaqus/CAE version 6.14-2 (2014). Dassault Systèmes Simulia Corp., Providence, RI, USA.

Simulia, 2014b. Abaqus/Explicit Version 6.14-2, Abaqus documentation . Dassault systemes, 2014.

van Staden, R.C., Guan, H., Johnson, N.W., Loo, Y., Meredith, N., 2008. Step-wise analysis of the dental implant insertion process using the finite element technique. Clin. Oral Implants Res. 19, 303–13. doi:10.1111/j.1600-0501.2007.01427.x